MEG Management Committee meeting

12.07.2010, 14:30hrs, West Wing Seminar Room, CBU

Minutes

- Present:  Friedemann Pulvermuller (Chair), Danny Mitchell, Yury Shtyrov, Olaf Hauk, Maarten Van Casteren, Robert Anthony, Georgie Morrill (Minutes).

- Apologies for absences

- Gary Chandler

- Rik Henson

- Clare Cook

- Lucy McGregor

- Matters arising

None.

- Minutes of the last meeting

Approved.

- Computing

MvC brought up the issue with saving MEG data.  There have been a few instances where it has been very slow and there has been concern over losing data.  RA reported that it was to do with disc space and this should be sorted now.  RA also reported that he hadn’t moved the database yet but this would be happening very soon.

It was agreed that RA/YS would wait for a Maxfilter update (probably December) before installing the new version.

MvC also asked if there was a way to access the on-line room scheduler remotely.  RA said that this was possible through running a VNC session.

- Lab business and progress

  - Eye tracker update

MvC reported that UCL are ready to test their new eye tracker again so OH and MvC will go and check this out.  Delays between the blink of the eye and where it appears on the recording should decrease.  Currently delays are around 45ms and with a new eye tracker, should decrease to around 1.5ms.

Action: MvC and OH visit UCL.

  - Acoustic door update

GC is still looking into this.

Action: GC look into acoustic door.

  - Other issues

OH confirmed that noise recordings are taking place but not analysed and YS/FP suggested uploading the data to a wiki page so all can see, then print and compare. OH also reported that UCL appear to perform regular phantom measurements and suggested doing this weekly at the CBU but it was considered unnecessary given the daily noise measurements, regular tuning routine (neither is available at UCL) and HPI fitting procedure as quality control measures.

Action: OH to improve the diagnostic scripts to make them more user-friendly. A possibility of keeping an archive of noise levels as web-pages in the lab’s site should be considered.

- Operator status approval: Lucy MacGregor

The committee gave approval for Lucy to be an operator for the MEG lab.

- MEG lab stats for June

In June, the MEG lab was operational 102 hours, with 146 hours available for experimental studies amounting to 70 percent used time.

- AOB

FP suggested having just 3 slots per day available in the MEG lab over July August, due to operators and many researchers being away in this period, making the lab not so in demand.  The committee agreed.

Action: GM to block out one slot per day on the resource scheduler.

Date of the next meeting

Monday 13th September at 2:30pm, in the WWSR.